{"text_format":"2","workbook":"xlsx","n_sheets":13,"sheets":["自动测试项总表","人工测试项总表","自动测试记录表","输入电压测试表","输入电流测试表","输入纹波测试表","输入浪涌测试表","转换效率测试表","输出电压测试表","输出电流测试表","输入输出延时测试表","输出纹波测试表","输出过冲测试表"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"自动测试项总表","index":0,"n_rows":75,"n_cols":28,"rows_kept":75,"cut":false,"hidden":false,"merged":[[56,0,56,3]]}
["设备配置表"]
["设备名称","设备类型","设备型号","控制句柄"]
["示波器1","示波器","DSO-X 3024A","GPIB0::11::INSTR"]
["示波器2","示波器","DSO-X 3024A","GPIB0::12::INSTR"]
["示波器3","示波器","DSO-X 3024A","GPIB0::13::INSTR"]
["示波器4","示波器","DSO-X 3024A","GPIB0::14::INSTR"]
["示波器5","示波器","DSO-X 3024A","GPIB0::15::INSTR"]
["示波器6","示波器","DSO-X 3024A","GPIB0::16::INSTR"]
["示波器7","示波器","DSO-X 3024A","GPIB0::17::INSTR"]
["示波器8","示波器","DSO-X 3024A","GPIB0::18::INSTR"]
["示波器9","示波器","DSO-X 3024A","GPIB0::19::INSTR"]
["示波器10","示波器","DSO-X 3024A","GPIB0::20::INSTR"]
["示波器11","示波器","DSO-X 3024A","GPIB0::21::INSTR"]
["示波器12","示波器","DSO-X 3024A","GPIB0::22::INSTR"]
["示波器13","示波器","DSO-X 3024A","GPIB0::23::INSTR"]
["示波器14","示波器","DSO-X 3024A","GPIB0::24::INSTR"]
["示波器15","示波器","DSO-X 3024A","GPIB0::25::INSTR"]
["示波器16","示波器","DSO-X 3024A","GPIB0::26::INSTR"]
["示波器17","示波器","DSO-X 3024A","GPIB0::27::INSTR"]
["示波器18","示波器","DSO-X 3024A","GPIB0::28::INSTR"]
["示波器19","示波器","DSO-X 3024A","GPIB0::29::INSTR"]
["示波器20","示波器","DSO-X 3024A","GPIB0::30::INSTR"]
["示波器21","示波器","DSO-X 3024A","GPIB0::31::INSTR"]
["示波器22","示波器","DSO-X 3024A","GPIB0::32::INSTR"]
["示波器23","示波器","DSO-X 3024A","GPIB0::33::INSTR"]
["示波器24","示波器","DSO-X 3024A","GPIB0::34::INSTR"]
["示波器25","示波器","DSO-X 3024A","GPIB0::35::INSTR"]
["示波器26","示波器","DSO-X 3024A","GPIB0::36::INSTR"]
["示波器27","示波器","DSO-X 3024A","GPIB0::37::INSTR"]
["示波器28","示波器","DSO-X 3024A","GPIB0::38::INSTR"]
["示波器29","示波器","DSO-X 3024A","GPIB0::39::INSTR"]
["示波器30","示波器","DSO-X 3024A","GPIB0::40::INSTR"]
["示波器31","示波器","DSO-X 3024A","GPIB0::41::INSTR"]
["示波器32","示波器","DSO-X 3024A","GPIB0::42::INSTR"]
["示波器33","示波器","DSO-X 3024A","GPIB0::43::INSTR"]
["示波器34","示波器","DSO-X 3024A","GPIB0::44::INSTR"]
["示波器35","示波器","DSO-X 3024A","GPIB0::45::INSTR"]
["示波器36","示波器","DSO-X 3024A","GPIB0::46::INSTR"]
["电子负载1","电子负载","IT8811","GPIB0::55::INSTR"]
["电子负载2","电子负载","IT8811","GPIB0::56::INSTR"]
["电子负载3","电子负载","IT8811","GPIB0::57::INSTR"]
["电子负载4","电子负载","IT8811","GPIB0::58::INSTR"]
["电子负载5","电子负载","IT8811","GPIB0::59::INSTR"]
["电子负载6","电子负载","IT8811","GPIB0::60::INSTR"]
["电子负载7","电子负载","IT8811","GPIB0::61::INSTR"]
["电子负载8","电子负载","IT8811","GPIB0::62::INSTR"]
["电子负载9","电子负载","IT8811","GPIB0::63::INSTR"]
["电子负载10","电子负载","IT8811","GPIB0::64::INSTR"]
["电子负载11","电子负载","IT8811","GPIB0::65::INSTR"]
["电子负载12","电子负载","IT8811","GPIB0::66::INSTR"]
["电子负载13","电子负载","IT8811","GPIB0::67::INSTR"]
["电子负载14","电子负载","IT8811","GPIB0::68::INSTR"]
["电子负载15","电子负载","IT8811","GPIB0::69::INSTR"]
["程控电源1","程控电源","IT6722","GPIB0::70::INSTR"]
[]
["测试指标配置表",{"$nulls":5},"+15V1",null,"-15V1",null,"+5V1",null,"-5V",null,"+12V",null,"+15V2",null,"-15V2",null,"+5V2",null,"+50V",null,"+25V",null,"+28V"]
["通用配置项",null,null,null,"输入",null,"CH1",null,"CH2",null,"CH3",null,"CH4",null,"CH5",null,"CH6",null,"CH7",null,"CH8",null,"CH9",null,"CH10",null,"CH11"]
["测试指标","默认勾选","测试模板","测试顺序","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态","测试参数配置表","启用状态"]
["输入电压","是","输入电压模板","1","输入电压测试表","启用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用","输入电压测试表","禁用"]
["输入电流","是","输入电流模板","2","输入电流测试表","启用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用","输入电流测试表","禁用"]
["输入纹波","是","输入纹波模板","3","输入纹波测试表","启用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用","输入纹波测试表","禁用"]
["输入浪涌","是","输入浪涌模板","4","输入浪涌测试表","启用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用","输入浪涌测试表","禁用"]
["输出电压","是","输出电压模板","5","输出电压测试表","禁用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用","输出电压测试表","启用"]
["输出电流","是","输出电流模板","6","输出电流测试表","禁用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用","输出电流测试表","启用"]
["输出纹波","是","输出纹波模板","7","输出纹波测试表","禁用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用","输出纹波测试表","启用"]
["输出过冲","是","输出过冲模板","8","输出过冲测试表","禁用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用","输出过冲测试表","启用"]
["输入输出延时","是","输入输出延时模板","9","输入输出延时测试表","禁用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用","输入输出延时测试表","启用"]
["转换效率","是","转换效率模板","10","转换效率测试表","启用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用","转换效率测试表","禁用"]
[]
[]
["公共参数表"]
["名称","类型","值","备注"]
["额定输入电压","double","28","单位：V"]
["下拉输入电压","double","19.0","单位：V"]
["上拉输入电压","double","36.0","单位：V"]

{"sheet":"人工测试项总表","index":1,"n_rows":5,"n_cols":4,"rows_kept":5,"cut":false,"hidden":false,"merged":[]}
["人工测试项总表"]
["测试内容","技术要求","测试值/结果","测试结论"]
["功耗测试","<10W"]
["温度老练","老练后加电正常"]
["工作温度","<65℃"]

{"sheet":"自动测试记录表","index":2,"n_rows":13,"n_cols":12,"rows_kept":13,"cut":false,"hidden":false,"merged":[]}
["自动测试记录表"]
[0,1,2,3,4,5,6,7,8,9,10,11]
["测试通道:","通道1","型号:","DC16XK","产品编号:",null,"检验员:"]
["设定电压:","28.00V","负载:","空载","循环次数:",null,"测试温度:",null,"相对湿度:",null,"测试结果:"]
["测试日期:","20210401","测试时间:","13:12","备注:"]
["输入电压:",null,"输入电流:",null,"输入浪涌:",null,"输入纹波:",null,"效率:"]
["Channel NO.","ch1","ch2","ch3","ch4","ch5","ch6","ch7","ch8","ch9","ch10","ch11"]
["输出通道:","+15V1","-15V1","+5V1","-5V","+12V","+15V2","-15V2","+5V2","+50V","+25V","+28V"]
["输出电压:"]
["输出电流:"]
["输出过冲:"]
["输出纹波:"]
["输入输出延时:"]

{"sheet":"输入电压测试表","index":3,"n_rows":13,"n_cols":5,"rows_kept":13,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["板卡通道","int","0","单位：ch"]
["额定阈值","double[]","27.0,29.0","单位：V"]
["下拉偏阈值","double[]","18.0,19.0","单位：V"]
["上拉偏阈值","double[]","36.0,37.0","单位：V"]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["额定输入电压","28",null,"28.1","27.9"]
["下拉偏输入电压","19",null,"19","18.3"]
["上拉偏输入电压","36",null,"36.3","36"]

{"sheet":"输入电流测试表","index":4,"n_rows":12,"n_cols":5,"rows_kept":12,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["板卡通道","int","1","单位：ch"]
["额定阈值","double[]","3.0,5.0","单位：A"]
["下拉偏阈值","double[]","5.0,6.0","单位：A"]
["上拉偏阈值","double[]","2.0,3.0","单位：A"]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]

{"sheet":"输入纹波测试表","index":5,"n_rows":17,"n_cols":5,"rows_kept":17,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["DSO通道","int[]","1,1","单位：ch"]
["DSO耦合","string","AC"]
["DSO通道SCALe","double","0.02","单位：V"]
["DSO通道偏移","double","0","单位：V"]
["DSOTIMeSCALe","double","0.01","单位：s"]
["DSOTIMe偏移","double","0","单位：s"]
["阈值","int[]","0,50","单位：mV"]
["通道延迟","int",250,"单位：ms"]
[]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]

{"sheet":"输入浪涌测试表","index":6,"n_rows":17,"n_cols":5,"rows_kept":17,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["DSO通道","int[]","1,1","单位：ch"]
["DSO耦合","string","DC"]
["DSO通道SCALe","double","0.02","单位：V"]
["DSO通道偏移","double","0","单位：V"]
["DSOTIMeSCALe","double","0.01","单位：s"]
["DSOTIMe偏移","double","0","单位：s"]
["阈值","int[]","0,50","单位：A"]
["通道延迟","int",250,"单位：ms"]
[]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]

{"sheet":"转换效率测试表","index":7,"n_rows":11,"n_cols":5,"rows_kept":11,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["阈值","int[]","70,100","单位：%"]
["通道延迟","int",250,"单位：ms"]
[]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]

{"sheet":"输出电压测试表","index":8,"n_rows":41,"n_cols":5,"rows_kept":41,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["+15V1空载阈值","double[]","15.0,15.1","单位：V"]
["-15V1空载阈值","double[]","-15.1,-15.0","单位：V"]
["+5V1空载阈值","double[]","4.9,5.1","单位：V"]
["-5V空载阈值","double[]","-5.1,4.9","单位：V"]
["+12V空载阈值","double[]","11.9,12.1","单位：V"]
["+15V2空载阈值","double[]","15.0,15.1","单位：V"]
["-15V2空载阈值","double[]","-15.1,-15.0","单位：V"]
["+5V2空载阈值","double[]","4.9,5.1","单位：V"]
["+50V空载阈值","double[]","49.5,50.5","单位：V"]
["+25V空载阈值","double[]","24.8,25.20","单位：V"]
["+28V空载阈值","double[]","27.5,28.5","单位：V"]
["+15V1半载阈值","double[]","14.9,15.1","单位：V"]
["-15V1半载阈值","double[]","-15.1,-14.9","单位：V"]
["+5V1半载阈值","double[]","4.9,5.1","单位：V"]
["-5V半载阈值","double[]","-5.1,4.9","单位：V"]
["+12V半载阈值","double[]","11.9,12.1","单位：V"]
["+15V2半载阈值","double[]","14.9,15.1","单位：V"]
["-15V2半载阈值","double[]","-15.1,-14.9","单位：V"]
["+5V2半载阈值","double[]","4.9,5.1","单位：V"]
["+50V半载阈值","double[]","49.5,50.5","单位：V"]
["+25V半载阈值","double[]","24.8,25.20","单位：V"]
["+28V半载阈值","double[]","27.5,28.5","单位：V"]
["+15V1满载阈值","double[]","14.9,15.1","单位：V"]
["-15V1满载阈值","double[]","-15.1,-14.9","单位：V"]
["+5V1满载阈值","double[]","4.9,5.1","单位：V"]
["-5V满载阈值","double[]","-5.1,4.9","单位：V"]
["+12V满载阈值","double[]","11.9,12.1","单位：V"]
["+15V2满载阈值","double[]","14.9,15.1","单位：V"]
["-15V2满载阈值","double[]","-15.1,-14.9","单位：V"]
["+5V2满载阈值","double[]","4.9,5.1","单位：V"]
["+50V满载阈值","double[]","49.5,50.5","单位：V"]
["+25V满载阈值","double[]","24.8,25.20","单位：V"]
["+28V满载阈值","double[]","27.5,28.5","单位：V"]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]

{"sheet":"输出电流测试表","index":9,"n_rows":41,"n_cols":5,"rows_kept":41,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["+15V1空载阈值","double[]","0,0.1","单位：A"]
["-15V1空载阈值","double[]","0,0.1","单位：A"]
["+5V1空载阈值","double[]","0,0.1","单位：A"]
["-5V空载阈值","double[]","0,0.1","单位：A"]
["+12V空载阈值","double[]","0,0.1","单位：A"]
["+15V2空载阈值","double[]","0,0.1","单位：A"]
["-15V2空载阈值","double[]","0,0.1","单位：A"]
["+5V2空载阈值","double[]","0,0.1","单位：A"]
["+50V空载阈值","double[]","0,0.1","单位：A"]
["+25V空载阈值","double[]","0,0.1","单位：A"]
["+28V空载阈值","double[]","0,0.1","单位：A"]
["+15V1半载阈值","double[]","1.09,1.11","单位：A"]
["-15V1半载阈值","double[]","0.29,0.31","单位：A"]
["+5V1半载阈值","double[]","3.19,3.21","单位：A"]
["-5V半载阈值","double[]","0.39,0.41","单位：A"]
["+12V半载阈值","double[]","0.69,0.71","单位：A"]
["+15V2半载阈值","double[]","0.34,0.36","单位：A"]
["-15V2半载阈值","double[]","0.39,0.41","单位：A"]
["+5V2半载阈值","double[]","3.19,3.21","单位：A"]
["+50V半载阈值","double[]","0.04,0.06","单位：A"]
["+25V半载阈值","double[]","0.49,0.51","单位：A"]
["+28V半载阈值","double[]","0,0.1","单位：A"]
["+15V1满载阈值","double[]","2.20,2.21","单位：A"]
["-15V1满载阈值","double[]","0.60,0.61","单位：A"]
["+5V1满载阈值","double[]","6.50,6.51","单位：A"]
["-5V满载阈值","double[]","0.80,0.81","单位：A"]
["+12V满载阈值","double[]","1.40,1.41","单位：A"]
["+15V2满载阈值","double[]","0.70,0.71","单位：A"]
["-15V2满载阈值","double[]","0.80,0.81","单位：A"]
["+5V2满载阈值","double[]","6.50,6.51","单位：A"]
["+50V满载阈值","double[]","0.1,0.11","单位：A"]
["+25V满载阈值","double[]","1.00,1.02","单位：A"]
["+28V满载阈值","double[]","0,0.1","单位：A"]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]

{"sheet":"输入输出延时测试表","index":10,"n_rows":19,"n_cols":5,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["+15V1阈值","double[]","0,30","单位：ms"]
["-15V1阈值","double[]","0,30","单位：ms"]
["+5V1阈值","double[]","0,50","单位：ms"]
["-5V阈值","double[]","0,50","单位：ms"]
["+12V阈值","double[]","0,45","单位：ms"]
["+15V2阈值","double[]","0,45","单位：ms"]
["-15V2阈值","double[]","0,50","单位：ms"]
["+5V2阈值","double[]","0,50","单位：ms"]
["+50V阈值","double[]","0,70","单位：ms"]
["+25V阈值","double[]","0,70","单位：ms"]
["+28V阈值","double[]","0,15","单位：ms"]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]

{"sheet":"输出纹波测试表","index":11,"n_rows":17,"n_cols":5,"rows_kept":17,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["DSO通道","int[]","1,1","单位：ch"]
["DSO耦合","string","AC"]
["DSO通道SCALe","double","0.02","单位：V"]
["DSO通道偏移","double","0","单位：V"]
["DSOTIMeSCALe","double","0.01","单位：s"]
["DSOTIMe偏移","double","0","单位：s"]
["阈值","int[]","0,50","单位：mV"]
["通道延迟","int",250,"单位：ms"]
[]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]

{"sheet":"输出过冲测试表","index":12,"n_rows":17,"n_cols":5,"rows_kept":17,"cut":false,"hidden":false,"merged":[]}
["参数表"]
["名称","类型","值","备注"]
["DSO通道","int[]","1,1","单位：ch"]
["DSO耦合","string","DC"]
["DSO通道SCALe","double","0.02","单位：V"]
["DSO通道偏移","double","0","单位：V"]
["DSOTIMeSCALe","double","0.01","单位：s"]
["DSOTIMe偏移","double","0","单位：s"]
["阈值","int[]","0,50","单位：V"]
["通道延迟","int",250,"单位：ms"]
[]
[]
[]
["测试显示表"]
["测试内容","技术要求","测试结果","上限值","下限值"]
["2GHz","-107±5",null,"-112","-102"]
["2.5GHz","-97±5",null,"-102","-92"]